## Cert
| Calibration Certificate |  |  |  | Certification Control: no. |  | 23196 | P
| Customer: | Siemens Energy - Sam Houston |  |  | Procedure ¦ Accuracy: |  | WS-WP-7.2
| Contact: | Steven Lumpkin |  |  | Calibration Date: |  | 2026-07-10
| Address: | 1200 West Sam Houston Parkway North |  |  | Due Date: |  | 2027-07-10
|  | Houston,TX 77043 |  |  | Frequency ¦ Interval: |  | 12 Months
| Phone: | 713-301-8678 |  |  | Temp (F) & Humidity: |  | 71.4° F ¦ 38%
| PO No.: |  |  |  | New/ Recal: |  | RECAL
| PSI | 30000 |  |  | S/N: | 2583
| Manuf. | MCDANIEL |  |  | Model: | ANALOG GAUGE
| As Found (Increasing) |  |  |  | As Left (Increasing)
| Pressure Gauge (PSI) | Standard Output (PSI) | Precent Deviation |  | Pressure Gauge (PSI) | Standard Output (PSI) | Precent Deviation
| 0 | 0 | 0% |  | 0 | 0 | 0%
| 6000 | 5973 | -0.45% |  | 6000 | 5973 | -0.45%
| 12000 | 11994 | -0.05% |  | 12000 | 11994 | -0.05%
| 18000 | 17983 | -0.09444% |  | 18000 | 17983 | -0.09444%
| 24000 | 23988 | -0.05% |  | 24000 | 23988 | -0.05%
| 30000 | 29971 | -0.09667% |  | 30000 | 29971 | -0.09667%
| As Found (decreasing) |  |  |  | As Left (decreasing)
| Pressure Gauge (PSI) | Standard Output (PSI) | Precent Deviation |  | Pressure Gauge (PSI) | Standard Output (PSI) | Precent Deviation
| 30000 | 29971 | -0.09667% |  | 30000 | 29971 | -0.09667%
| 24000 | 23988 | -0.05% |  | 24000 | 23988 | -0.05%
| 18000 | 17983 | -0.09444% |  | 18000 | 17983 | -0.09444%
| 12000 | 11994 | -0.05% |  | 12000 | 11994 | -0.05%
| 6000 | 5973 | -0.45% |  | 6000 | 5973 | -0.45%
| 0 | 0 | 0% |  | 0 | 0 | 0%
Calibration Standard
Brand: | Item: | Model | S/N | Due Date |  | Traceability |  | Accuracy
ADDITEL | GAUGE | ADT-681-60K | 218231A0029 | 2027-06-16 |  | N24061747 |  | ±0.25%
The calibrations within the certificate/report are traceable through NIST or another National Metrology Institute to the International System of Units (SI). The results only relate to the items tested, calibrated, or sampled. The reported expanded uncertainty of measurement is stated as the standard uncertainty of measurement multiplied by the coverage factor k=2, which for a normal distribution corresponds to a coverage probability of approximately 95%. Simple acceptance decision rule were uncertainty is not factored into the tolerance for conformity.
Calibrated by: |  | David Bain |  |  | Approved by: |  | DAVID SHAW
|  | E-Signed: | 2026-07-10 |  |  |  | E-Signed: | 2026-07-10
Comments:
Formulas:
Due Date = VLOOKUP([Siemens Energy - Sam Houston],'Data - Standards'!A2:I18,5,FALSE)
P = [Precent Deviation]

## DATA - CUSTOMER
<Choose Company> | Company | Contact | Addr1 | Addr2 | City, State,Zip |  | Zip | Country | Phone | Fax | blah | blah | blah | Old
3C METAL | 3C Metal | Edourd Villaret | 11983-A FM 529 |  | Houston, TX 77056 | TX | 77041 | USA | (713) 808-9651 |  |  |  |  | 2
3D At Depth | 3D At Depth | Alfredo Escobedo | 2901 W Sam Houston Pkwy N Suite A-150 |  | Houston, TX 77043 | TX | 77043 | USA | (713) 966-6779
A&M Equipment LLC. | A&M Equipment LLC. | Keith Matetzschk | 28246 FM 2920 Rd. |  | Waller, TX 77484 | Tx | 77484 | USA | (979) 525-4431
ACE WINCHES | ACE WINCHES | STEVEN SMITH | TOWIE BARCLAY WORKS |  | TURRIFF, ABERDEENSHIRE |  | AB538EN | UK | 44 7768 067324
ACFM | Howden | Roosevelt Bruins | 16250 Tomball Pkwy |  | Houston, TX 77086 | TX | 77086 | USA | (936) 539-5252 | (936) 539-5251 |  |  |  | 3
AIR PRODUCT | AIR PRODUCT | JOSE G CONTRERAS | 10202 STRANF RD |  | LAPORTE | TX |  | USA | 832-547-1897
Baker Hughes - Brittmore | Baker Hughes | James Brooks | 6500 Brittmore RD. |  | Houston, TX 77041 | TX | 77041 | USA | (281) 529-2400
Bestway Oilfield Inc. | Bestway Oilfield Inc. | Hector Lopez | 3471 W. State Hwy 44 |  | Alice, TX 78333 | Tx | 78333 | USA | 281-452-2525
Bolttech Mannings | Bolttech Mannings | Quality Department | 290 Beltway Green Blvd |  | Deer Park, TX 77536 | Tx | 77536 | USA
C-Automation | C-Automation | Randall Stockton | 10535 Fisher Rd |  | Houston, TX 77041 | TX | 77041 | USA | 832-467-4644
DRIL QUIP | Dril Quip | EFRAIN VILLASENOR | 6401 N Eldridge Pkwy |  | Houston, TX 77041 | Tx | 77041 | USA | 281-425-5086
D-Tech Calibration | D-Tech Calibration | Calibration Department | 16151 Cairnway Dr. Suite 100 |  | Houston, TX 77084 | TX | 77084 | USA | 832-674-4961
Enermech | Enermech | Xavier Ochoa | 12101 CUTTEN RD |  | HOUSTON, TX 77066 | TX | 77066 | USA | 979-571-2820 |  |  |  |  | 5
Engineered Custom Solutions, | Engineered Custom Solutions, LLC. | Tam Nguyen | 13720 FM 529 Rd |  | Houston, TX 77041 | TX | 77041 | USA | (832) 598-2083
Fastenal | Fastenal | Jacob Tackett | 201 Bryan St. |  | Dayton, TX 77535 | Texas | 77535 | USA | (936) 258-2671 |  |  |  |  | 6
Fastenal (#2) | Fastenal | Chris Conrad | 501B East Heil Ave. |  | El Centro, CA 92243 | CA | 92243 | USA | (760)-352-2313 |  |  |  |  | 7
Fastenal Clute | Fastenal Clute | Rodney Kifer | 734 S Main St, |  | Clute, TX 77531 | tx | 77531 | USA | 979-265-2252
Flatiron | Flatiron | Micah Murphy | 817 Brewster ST |  | Corpus C, TX 78401 | Tx | 78401 | USA | 281-946-4000
Genesis Solar | Genesis Solar | Michael Medsker | 11995 Wiley's Well Rd |  | Blythe, CA 92225 | CA | 92225 | USA | 760-921-1409
GRIZZLY | Grizzly | Kirk French | 41303 park 290 drive |  | Waller, TX 77484 | Texas | 77484 | USA | 710-926-5544 |  |  |  |  | 8
Herc Rentals Inc. | Herc Rentals Inc. | Lin Mendez | 1002 Clay Ct |  | Deer Park, TX 77536 | TX | 77536 | USA | 281-478-7921
HOUSTON PRECISION | HOUSTON PRECISION | ANDREA DeLEON | 6633 POLK STREET |  | Houston, TX 77011 | TX | 77011 | USA | [5 empty] | 27
HP Piping | HP Piping | Thomas Mitchmore | 4311 holmes Rd |  | Houston, TX 77051 | Texas | 77051 | USA | 832-582-8898 |  |  |  |  | 9
Hughes-Anderson Heat Exchanger | Hughes-Anderson Heat Exchanger | Kelly Bock | PO Box 582710 |  | Tulsa, OK 74158 | OK | 74158 | USA | 918 836-1681
Hurricane Tool | Hurricane Tool | Brittany Mason | 2330 Pasadena Blvd., Bldg A |  | Pasadena, TX 77502 | TX | 77502 | USA | 281-476-5530
HYTORC | HYTORC
Hytorc North | Hytorc North | Tina Moss | 311 Taft St |  | Palmyra, WI 53156 | WI | 53156 | USA | 262-510-9587
INTEGRA Technologies | INTEGRA Technologies | Jared Priddy | 5812 Genoa Red Bluff Rd |  | Pasadena, TX 77507 | Tx | 77507 | USA | 713-920-2400
K2 | Applus K2 | James Rushforth | 11801 West Sam Houston Parkway |  | Houston, TX 77031 | TX | 77031 | USA | (832) 454-1555 |  |  |  |  | 10
Kirksey Machine | Kirksey Machine | Cody Moore | 8721 Clay Rd. |  | Houston, TX 77292 | Tx | 77292 | USA | 713-682-6704
MAN Energy Solutions USA Inc. | MAN Energy Solutions USA Inc. | Curtis Chandler | 1758 Twinwood Pkwy |  | Brookshire,Tx 77423 | Tx | 77423 | USA | (713) 780-4200
MSI | Measurement Systems and Services, Inc. | David Klasen | 5010 Steffani Lane |  | Houston, TX 77041 | TX | 77041 | USA | - |  |  |  |  | 11
Mustang CAT | Mustang CAT | Alfred James III | 12800 Northwest Freeway |  | Houston, TX 77040 | TX | 77040 | USA | 713-452-7205 |  |  |  |  | 27
Nabors | Nabors Drilling Solutions | Joey Peregne | 515 W. Greens Road Suite 110 |  | Houston, TX 77067 | TX | 77067 | USA | (281) 874-0035 |  |  |  |  | 12
Nabors Breen | Nabors Breen | Jeremy X | 8400 Breen Rd |  | Houston, Tx 77064 | Tx | 77064 | USA | 832-405-8697
NABORS CROSBY | NABORS CROSBY | Orlando Sanchez | 2100 Crosby Dayton Rd |  | Crosby, TX 77532 | TX | 77532 | USA | 281.618.4111
Nextera Energy | Nextera Energy | Michael Medsker | 601 Travis St |  | Houston, 77002 | TX | 77002 | USA | 760-921-1409 |  |  |  |  | 13
NOV | Metrology Lab WLY/NOV | Roosevelt Perez | 12950 West Little York Rd |  | Houston, TX 77041 | Texas | 77041 | USA | 713-501-7784 |  |  |  |  | 14
NOV 529 | NOV/ Metrology Lab | Eric Aguirre | 11919 FM 529 RD |  | Houston, TX 77041 | TX | 77041 | USA | 832-424-7774 |  |  |  |  | 15
NOV CALEXICO | NOV CALEXICO | ANDREA LEDESMA | 2270 PORTICO BLVD |  | Calexico, CA 92231 | CA | 92231 | USA | - |  |  |  |  | 16
NOV CREEKMOUNT | NOV CREEKMOUNT | Luis Gonzalez | 1200 W LITTLE YORK RD |  | Houston, TX 77091 | Tx | 77091 | USA | (713)-957-2288
NOV Rubber Plant | Metrology Lab WLY/NOV | Roosevelt Perez | 12950 West Little York Rd |  | Houston, TX 77041 | Texas | 77041 | USA | 713-501-7784 |  |  |  |  | 14
NOV Top Drive | NOV Top Drive | Luis Montes | 12950 W. Little York |  | Houston, TX 77041 | TX | 77041 |  | -
OES Oilfield Services | OES Oilfield Services | Quality Department | 8615 Jackrabbit Rd |  | Houston, TX 77095 | TX | 77095 | USA | (713) 960-1339 |  |  |  |  | 27
OGME | OGME | ROBERT AITKEN | 5227 DOW ROAD |  | Houston, TX 77040 | TX | 77040 | USA | 713-686-5783 |  |  |  |  | 17
Opero Energy | Opero Energy | Evan Marthinson | 12261 FM 529 |  | Houston, TX 77041 | Tx | 77041 | USA | 713-452-3229
OSIES | OSIES INC. | Jennifer Alvarez | 12734 Tanner Rd |  | Houston, TX 77041 | TX | 77041 | USA | (713) 849-5131 | (713) 849-5148 |  |  |  | 18
PetrolValves Group | PetrolValves Group | Dennis Perez | 6000 W. Sam Houston Parkway N. |  | Houston, TX 77041 | Tx | 77041 | USA | (713) 856-0404
Pivot Engineers | Pivot Engineers | Quality Department | 1114 Lost Creek Blvd |  | Austin, TX 78746 | TX | 78746 | USA | 512-893-3947
Plumbaway Plumbing | Plumbaway Plumbing | Robert Catano | 122 Floradora Ln |  | Houston, TX 77076 | TX | 77076 | USA | 832-5478799
Power Tool Service | Power Tool Service | JB Robertson | 3718 Polk Street |  | Houston, TX 77003 | TX | 77003 | USA | 713-741-8127
Precision Drilling | Precision Drilling | Alex Ventura | 14041 Vickery Dr. |  | Houston,TX 77032 | TX | 77032 | USA | 281-227-5753
Reytec Construction Resources, Inc | Reytec Construction Resources, Inc | Sophie Venturini | 1901 Hollister Street |  | Houston, TX 77080 | Tx | 77080 | USA | 832.212.9354
SGRE | SIEMENS GAMESA - Lower Snake River Wind | Tina Moss | 39 Falling Springs Road |  | POMEROY, WA 99347 | TX | 99347 | USA | 262.495.2656
Ship Channel Constructors LLC | Ship Channel Constructors LLC | Jeff Miron | 15015 East Freeway B |  | Channelview, TX 77530 | TX | 77530 | USA | 361-677-1091
SIEMENS ENERGY | SIEMENS ENERGY | BILLY EBERT | 405 DEERWOOD GLEN |  | Deer Park, TX 77536 | TX | 77536 | USA | 281-946-4000 |  |  |  |  | 19
Siemens Energy - Sam Houston | Siemens Energy - Sam Houston | Steven Lumpkin | 1200 West Sam Houston Parkway North |  | Houston,TX 77043 | TX | 77043 | USA | 713-301-8678
SIEMENS GAMESA - Fairless Hills | SIEMENS GAMESA - Fairless Hills | Bobby Kuzemka | 290 Canal Road |  | Fairless Hills, PA 19030 | PA | 19030 | USA | 267-793-5426 |  |  |  |  | 19
Siemens Gamesa - Bison Wind Farm | Siemens Gamesa - Bison Wind Farm | Tiffany Mahoney | 5198 30th street |  | New Salem, ND 58563 | ND | 58563 | USA | -
Siemens Gamesa- Abilene | Siemens Gamesa- Abilene | Roland Carrington | 4400 Alafaya Trail |  | Orlandox FL 32826 | TX | 32826 | USA | (361) 222-4217
Siemens Gamesa- OKC | Siemens Gamesa- OKC | Tiffany Mahoney | 7501 SW 29th Street |  | Oklahoma City, OK 73179 | OK | 73179 | USA | 689-220-9110
SIEMENS IOWA | Siemens Gamesa Renewable Energy Inc. | Jack Lavoy | 106 East Hillcrest Street |  | Adair, IA 50002 | IA | 50002 | USA | -
SIEMENS WTC | SIEMENS WTC | Melaney Glendening | 1090 E. 37th St. N. |  | Wichita, KS 67219 | TX | 67219 | USA | 316-315-4534
Sparrows Group | Sparrows Group | Garrett Lege | 6707 Northwinds Drive |  | Houston, Tx 77041 | Tx | 77041 | USA | 832-467-7300
Specialloy Instrumentation | Soecialloy Instrumentation | Chris Sneed | 11702 Brittmoore Park Dr. |  | Houston, TX 77041 | TX | 77041 | USA | 713-827-7776 |  |  |  |  | 20
Starcon | Starcon | Francisco Aguilar | 10610 W. Fairmont Parkway |  | La Porte, TX 77571 | TX | 77571 | USA | 281-291-5260
STATS | Stats Internatinal Inc | Daelynn Williams | 10950 Britmoore Park Dr. |  | Houston, TX 77041 | TX | 77041 | USA | (713) 896-3150 |  |  |  |  | 21
Subsea Americas Houston | Subsea Americas Houston | Charles | 1611 Elmview Drive Building A |  | Houston, TX 77080 | Tx | 77080 | USA | 985-714-1767
TFMC | TFMC | Kenneth Fox | 1805 Brittmoore Rd |  | Houston, TX 77043 | TX | 77043 | USA | 713-301-6695
Thrustmaster of Texas | Thrustmaster of Texas | Maria Lopez | 6900 Thrustmaster Drive |  | Houston, TX 77041 | TX | 77041 | USA | 713-937-6295 |  |  |  |  | 22
Tops Field Services LLC | Tops Field Services LLC | Mike Earnest | 3221 5th Ave S |  | Texas City, TX 77590 | TX | 77590 | USA | 281-764-5205
Tri-Torc | Tri-Torc INC. | Mr. Sandeep Shukla | 2006 Wichita St. |  | Pasadena, TX 77502 | TX | 77502 | USA | 832-831-4342 |  |  |  |  | 23
Turner Supply | Turner Supply | Mike Johnson | 405 Deerwood Glen Drive |  | Deer Park, TX 77536 | TX | 77536 | USA | N/A
Universal Plant Services | Universal Plant Services | James Buck Mooney | 4460 Highway 225 |  | Deer Park, TX 77536 | TX | 77536 | USA | 281-638-5663
USA DeBusk | USA De Busk | Marc Fernandez | 1101 w 8th st. |  | Deer Park, TX 77536 | TX | 77536 | USA | 281-979-5269 |  |  |  |  | 4
VOGT | VOGT Valves | David Bain | 13800 North Promenade |  | Stafford, TX 77477 | TX | 77477 | USA | 281-313-8080 |  |  |  |  | 24
Weldfit | Weldfit | Daniel DeSmet | 4133 Southerland Rd |  | Houston, TX 77092 | TX | 77092 | USA | 7134603700
Winshaw | Winshaw Hydraulic Tools | David Shaw | 6403 Brittmoore Road |  | Houston, TX 77041 | Tx | 77041 | USA | 281-885-8561 |  |  |  |  | 25
Winshaw Hydraulic Tools | Winshaw Hydraulic Tools(RENTAL) | Rental Coordinator | 6403 Brittmoore Road |  | Houston, TX 77041 | Tx | 77041 | USA | 281-885-8561 |  |  |  |  | 25
WOM | Worldwide Oilfield Machine, inc | Jorge A. B arron | 11400 Tanner Rd, |  | Houston, TX 77041 | TX | 77041 | USA | 832-547-1270 |  |  |  |  | 26
Yokohama TWS | Yokohama TWS | Ricardo Orozco | 1811 Brittmoore RD. suite 500 |  | Houston, TX 77043 | TX | 77043 | USA | 832-848-4949

## Data - Standards
Model | Brand: | Item: | S/N | Due Date | Traceability | Accuracy | Current Date | Days Left | Ownership
XP2i 15k | CRYSTAL | GAUGE | 562009 | 2028-10-09 | 1194717 | ±0.25% | 2026-07-13 | 819 | WINSHAW | Recall Interval: | New / Recal: | Testing Procedure:
XP2i 10k | CRYSTAL | GAUGE | 561154 | 2026-10-22 | A1389576 | ±0.25% | 2026-07-13 | 101 | WINSHAW | 3 Months | NEW | WS-WP-7.2
ADT-681-60K | ADDITEL | GAUGE | 218231A0029 | 2027-06-16 | N24061747 | ±0.25% | 2026-07-13 | 338 | WINSHAW | 6 Months | RECAL | CP-43 - Torque - Rev. 6
CPB5000 | bedenburg | GAUGE | 34342 | 2026-04-26 | 30644R |  | 2026-07-13 | -78 |  | 9 Months |  | CP-56 - Torque - Rev.6
2,500 - 60,000 LBF | COTI INC | GAUGE | 5-1017 | 2026-07-24 | 1141729 | ±1% | 2026-07-13 | 11 | Holloway Houston | 12 Months |  | CP-76 - Pressure - Rev.7
SW30-100K-B363 | TOVEY | Pressure Transducer | 132312A | 2028-02-26 | 684.07/O-0000058461 | ±1% | 2026-07-13 | 593 | WINSHAW | 24 Months |  | COMPRESSION
7230 | TOVEY | Digital Indicator | 17818127 | 2028-02-26 | 684.06/O-0000035085-21 | ±1% | 2026-07-13 | 593 | WINSHAW |  |  | COPS-QM-PRO-003 REV:F
[8 empty] | 0 |  |  |  | CP-50 - Weight - Rev. 4
30k Chart Record | Barton | 30k Chart Record | 242E-060413R | 2025-10-16 | 154022 | ±1% | 2026-07-13 | -270 |  |  |  | CP-100 - Fluid Control - Rev.2
THERMOMETER | TRACE | THERMOMETER | 200141856 | 2026-11-24 | 35499 | ±0.1% | 2026-07-13 | 134 | OGME